## Sheet1
ReqID | Clause | Title | Requirement (paraphrased) | Acceptance Criteria (evidence to look for) | Expected Artifacts / Records | Notes & Guidance (from ISO/TR 24971 where applicable)
ISO14971-4.1-01 | 4.1 | Risk management process established | Establish, implement, document, and maintain an ongoing risk management process covering the entire device life cycle. Include risk analysis, risk evaluation, risk control, and production & post‑production activities. Integrate with product realization as appropriate. | Procedure(s) exist and are controlled; scope includes life‑cycle coverage; process map/flow shows required elements; references to related QMS processes are clear. | Risk Management SOP; process flowchart; quality manual/QMS cross‑references. | TR24971 clarifies the process can be within a QMS but is not required; ensure linkage to ISO 13485 processes as relevant.
ISO14971-4.2-01 | 4.2 | Top management commitment | Provide evidence of top‑management commitment: adequate resources and assignment of competent personnel. | Documented management commitment (e.g., management review minutes, resourcing decisions); defined roles; training plans. | Management Review records; org charts; resource plans. | TR24971 details the emphasis on top management and periodic suitability review.
ISO14971-4.2-02 | 4.2 | Policy for risk acceptability | Define and document a policy for establishing criteria for risk acceptability, aligned with regulations, relevant standards, state of the art, and stakeholder concerns; include guidance on overall residual risk. | Approved policy exists; cites regulatory/standard inputs; defines principles for specific device families where needed; accessible to teams. | Risk Acceptability Policy document; references within RMPs. | TR24971 Annex C provides detailed guidance and examples for setting the policy and criteria.
ISO14971-4.2-03 | 4.2 | Suitability review of the process | Review the suitability and effectiveness of the risk management process at planned intervals; record decisions and actions. | Evidence of periodic reviews (inputs/outputs, action tracking); updates to procedures as needed. | Management Review records; CAPA/Change control tickets. | TR24971 lists aspects to review (procedure effectiveness, criteria adequacy, feedback loop effectiveness).
ISO14971-4.3-01 | 4.3 | Competence of personnel | Ensure persons performing risk management tasks are competent (education, training, skills, experience) and maintain objective evidence of competence. | Training matrix and records exist; CVs/experience captured; role‑based competency profiles defined. | Training records; job descriptions; competency matrices. | Annex A & TR24971 include example competencies by function; such records need not be inside the RMF.
ISO14971-4.4-01 | 4.4(a) | RMP scope & lifecycle coverage | RMP identifies the device and the life‑cycle phases for which each element applies. | RMP names device and versions; matrices mapping plan elements to life‑cycle phases. | Risk Management Plan (RMP) vX; lifecycle applicability table. | Annex A explains why scope is critical for objectivity.
ISO14971-4.4-02 | 4.4(b) | Responsibilities & authorities | Assign responsibilities and authorities for all risk management activities. | Named owners/approvers for each activity; RACI chart in plan. | RMP section with RACI; signatures/approvals. | TR24971 shows typical roles and knowledge areas.
ISO14971-4.4-03 | 4.4(c) | Review of RM activities | Define requirements for reviews of RM activities (e.g., design reviews, periodic RM check‑ins). | Planned reviews with frequency/triggers; meeting minutes stored; issues tracked to closure. | Design review minutes; RM review logs. | Align reviews with design controls and PMS feedback.
ISO14971-4.4-04 | 4.4(d) | Criteria for risk acceptability | Define criteria for risk acceptability, including cases where probability of occurrence of harm cannot be estimated. | Risk criteria table/matrix defined; rules for ‘probability unknown’ documented; applies per device scope. | Risk criteria appendix in RMP; risk matrix legend. | TR24971 offers qualitative/semi‑quantitative scales and cautions about matrix design & rationale.
ISO14971-4.4-05 | 4.4(e) | Method for overall residual risk | Define the method and criteria for evaluating overall residual risk, in relation to benefits of intended use. | Documented method (e.g., aggregation + expert/clinical judgement); criteria defined; sources of benefit evidence identified. | Overall Residual Risk Method section; clinical/benefit evidence list. | TR24971 gives examples of methods and evidence sources.
ISO14971-4.4-06 | 4.4(f) | Verification planning | Plan verification of implementation and effectiveness of risk control measures. | Verification strategy linked to each control; protocols defined or referenced; pass/fail criteria. | Verification protocol(s); V&V plan; test reports. | Annex A stresses early planning of verification resources.
ISO14971-4.4-07 | 4.4(g) | P&PP information collection & review | Define methods to collect and review production and post‑production information and feed back into RM. | Defined sources, cadence, ownership, analysis methods, escalation routes. | PMS plan; complaint handling SOP; literature search plan. | TR24971 aligns with ISO 13485 §§7–8 and ISO/TR 20416 PMS guidance.
ISO14971-4.4-08 | 4.4(h) | Plan change control | Maintain a record of changes to the RMP over the device life cycle. | Version history present; changes justified and approved; RMF updated. | RMP revision log; change control tickets. | Annex A notes change records facilitate audit and review.
ISO14971-4.5-01 | 4.5 | Risk Management File (RMF) & traceability | Establish and maintain an RMF. Provide traceability from each identified hazard to risk analysis, risk evaluation, implementation & verification of controls, and residual risk evaluation. | Traceability matrix complete for all hazards; RMF collects or points to all required records; retrievable in timely fashion. | RMF index; traceability matrix/baseline; pointers into QMS repositories. | TR24971 clarifies RMF is a logical construct; can point to other systems; recommends a hazard‑centric index.
ISO14971-5.1-01 | 5.1 | Conduct & record risk analysis | Perform risk analysis per 5.2–5.5 and record results, including device ID, analysts/organizations, scope, and date. | Analysis package includes required identifiers and dates; results trace to RMF. | Risk Analysis dossier; cover sheet with meta‑data. | TR24971 provides technique guidance (e.g., FMEA, FTA); reuse of similar device analyses allowed with justification.
ISO14971-5.2-01 | 5.2 | Intended use documented | Document intended use considering indication, patient population, body part/tissue, user profile, use environment, and operating principle. | Use specification present and consistent across DHF and labeling; stored in RMF. | Use specification; user needs; intended purpose statement. | TR24971 lists factors and links to usability engineering inputs.
ISO14971-5.2-02 | 5.2 | Reasonably foreseeable misuse | Document reasonably foreseeable misuse; consider predictable human behaviour and use errors. | Misuse scenarios identified and justified; linkage to HFE/validation or field data. | Misuse log; HFE reports; post‑market signals. | TR24971 explains deriving misuse via simulated/actual use and HFE.
ISO14971-5.3-01 | 5.3 | Characteristics related to safety | Identify and document qualitative/quantitative characteristics that could affect safety; define limits where appropriate. | Characteristics list complete with limits/tolerances; link to essential performance where applicable. | Characteristics inventory; engineering specs; acceptance limits. | TR24971 offers guiding questions to elicit characteristics.
ISO14971-5.4-01 | 5.4 | Hazards & hazardous situations | Identify known and foreseeable hazards; for each hazard, identify sequences/combinations of events leading to hazardous situations across life‑cycle stages. | Comprehensive hazard list with linked hazardous situations; life‑cycle coverage; interfaces considered. | Hazard list; event trees/FTAs; interface analysis. | TR24971 Annexes give examples and techniques; systematic methods recommended.
ISO14971-5.5-01 | 5.5 | Risk estimation per hazardous situation | For each hazardous situation, estimate the associated risk(s) using available data; define severity and probability approach; handle cases where probability cannot be estimated per policy. | Documented rationale, data sources, and calculation/scoring; treatment for ‘unknown probability’ aligns with RMP. | Risk estimation tables; datasets; literature review logs. | TR24971 provides example scales and matrices; document rationale and avoid over‑granularity.
ISO14971-6-01 | 6 | Risk evaluation vs criteria | Compare estimated risks against criteria for acceptability; if acceptable, treat as residual risk; otherwise proceed to risk control. | Evaluation outcome recorded for each hazardous situation; consistent application of criteria. | Risk evaluation tables; sign‑offs. | TR24971: manufacturer sets criteria in policy/RMP; examples in Annex C.
ISO14971-7.1-01 | 7.1 | Risk control option analysis (priority order) | Determine risk control measures and select options in priority: (a) inherently safe design/manufacture; (b) protective measures; (c) information for safety and training. | Controls chosen reflect priority order unless justified; rationale documented; relevant standards considered. | Risk control option analysis log; design changes; protective features; IFU/training plans. | TR24971 expands examples for each option and their effect on probability/severity.
ISO14971-7.1-02 | 7.1 | Record selected measures & impracticability | Record selected measures; if further reduction is not practicable, trigger benefit‑risk analysis. | Records show selected controls and any impracticability analysis with justification. | RCA for impracticability; technical assessments. | Align with state of the art and feasibility analysis.
ISO14971-7.2-01 | 7.2 | Implement selected measures | Implement all selected risk control measures. | Links from controls to design outputs/process controls; implementation evidence stored. | Design outputs; process qualifications. | —
ISO14971-7.2-02 | 7.2 | Verify implementation | Verify implementation of each risk control measure. | Objective evidence that each control exists as designed (tests/inspection). | Verification protocols & reports; DHR evidence. | Verification may occur in design verification or process qualification.
ISO14971-7.2-03 | 7.2 | Verify effectiveness | Verify the effectiveness of risk control measures in reducing risk. | Evidence that severity/probability reduction was realized; method appropriate (validation, user testing, etc.). | Validation reports; user studies; performance tests. | Examples include dose accuracy verification or process qualification linking to risk reduction.
ISO14971-7.3-01 | 7.3 | Residual risk evaluation (per HS) | After implementing controls, evaluate residual risk for each hazardous situation against acceptability criteria; iterate if not acceptable. | Residual risk recorded; next actions defined for non‑acceptable risks. | Residual risk tables; decision logs. | —
ISO14971-7.4-01 | 7.4 | Benefit‑risk analysis (BRA) | If residual risk is not acceptable and further control is not practicable, perform BRA to determine if benefits outweigh residual risk. | BRA present with data/literature; conclusion documented; if benefits do not outweigh risk, modify device or intended use. | BRA report; clinical/market data; expert panel minutes. | Follow TR24971 guidance for BRA execution and documentation.
ISO14971-7.5-01 | 7.5 | Risks arising from controls | Review whether risk controls introduce new hazards or affect previously estimated risks; manage any new/increased risks. | Assessment documented post‑implementation; new risks fed back into analysis. | Change impact assessments; updated analyses. | —
ISO14971-7.6-01 | 7.6 | Completeness of risk control | Review risk control activities for completeness across all identified hazardous situations. | Checklist completed; evidence that all planned controls were implemented and verified. | Completeness review record; sign‑off. | —
ISO14971-8-01 | 8 | Overall residual risk evaluation | After implementing and verifying all controls, evaluate overall residual risk in relation to benefits, using method/criteria defined in RMP. | Documented overall evaluation with conclusion acceptable/not; references to supporting evidence. | Overall residual risk evaluation memo; clinical benefit dossier. | TR24971 provides guidance on methods and disclosure of significant residual risks.
ISO14971-8-02 | 8 | Disclosure of significant residual risks | Inform users of significant residual risks in accompanying documentation. | Labeling/IFU includes residual risk disclosures; cross‑referenced in RMF. | IFU, labeling change history; RMF cross‑refs. | —
ISO14971-9-01 | 9 | Risk management review & report | Prior to commercial release, review execution of the RMP to ensure plan was implemented, overall residual risk is acceptable, and P&PP methods are in place; record results as the Risk Management Report (RMR) and include in RMF; responsibility assigned in RMP. | Approved RMR present; confirms criteria; identifies any conditions of release; reviewer has appropriate authority. | Risk Management Report; RMP section listing responsible reviewer; approvals. | TR24971 notes RMR can be updated when new information emerges.
ISO14971-10.1-01 | 10.1 | System for P&PP monitoring | Establish, document, and maintain a system to actively collect and review safety‑relevant information during production and post‑production; consider appropriate collection and processing methods. | PMS system/procedures defined; active and passive sources listed; responsibilities and tools defined. | PMS Plan; complaint handling SOP; vigilance/field action procedures; data pipelines. | TR24971 emphasizes the feedback loop to keep RM continuous; aligns with ISO 13485 and ISO/TR 20416.
ISO14971-10.2-01 | 10.2 | Information collection — sources | Collect information as applicable: production/process monitoring; user and servicing feedback; supply chain inputs; publicly available information; and state of the art; also consider similar devices and other products. | Evidence of routine ingestion from listed sources; periodic literature/signal review; competitor/alternative therapy monitoring. | Supplier quality records; service logs; complaint database; literature review logs. | TR24971 explains typical sources and the value of prior experience with similar devices.
ISO14971-10.3-01 | 10.3 | Information review & escalation | Review collected information for relevance to safety (e.g., new hazards, risks no longer acceptable) and feed back into earlier RM phases with appropriate actions. | Signal review minutes; triggers documented; updates to hazard analysis/RMP; CAPA as needed. | Signal management SOP; safety review board minutes; RM updates. | TR24971 details the continuous feedback loop and alignment with PMS.
TR24971-AnnC-01 | TR §6 & Annex C | Risk matrices & scales documented | If using qualitative or semi‑quantitative scales/matrices, document definitions, ranges, and rationale for chosen matrix size and cut‑points. | Matrix legend and definitions approved; justification for scale choice; avoid excessive levels without data support. | Risk matrix appendix; probability/severity definitions. | TR24971 includes example 3×3 and 5×5 matrices and warns about data needs for >5 levels.
TR24971-Trace-01 | TR §4.5 | Traceability approach documented | Document the traceability method/tool linking hazards → analysis → evaluation → controls → residual risk → overall evaluation. | Traceability index is complete and stays current through changes; retrievable quickly. | Traceability matrix; RMF index. | TR24971 recommends a hazard‑centric index and keeping it current as devices change.