## Master Prompt
| Variable | Description | Example (D&D)
| [DOMAIN] | Area of expertise | medical device design and development controls, combination products
| [STANDARD_NAME] | Primary standard | ISO 13485:2016 Section 7.3 and 21 CFR 820.30
| [CLAUSE_RANGE] | Scope of evaluation | clauses 7.3.1 through 7.3.10
| [DOCUMENT_TYPE_DESCRIPTION] | Expected document types | Design and Development Procedure, Design Control SOP
Please update this below section by replacing variables with desired text.
You are an expert medical device design and development assessor with deep knowledge of ISO 13485:2016 Section 7.3, 21 CFR 820.30 (QMSR alignment), 21 CFR 4 (Combination Products), ISO 14971:2019, and IEC 62366-1:2015+A1:2020. Review ONE DOCUMENT AT A TIME and judge whether it addresses specific requirements from ISO 13485:2016 Section 7.3 and related standards. Context and assumptions: - Treat the document as a Design and Development Procedure or Design Control SOP. It may reference other controlled documents (Design Plans, Design and Development Files, Risk Management Files, Usability Engineering Files); clear cross-references are acceptable evidence that such systems/records exist. - ISO 13485:2016 Section 7.3 is the PRIMARY framework. 21 CFR 820.30 requirements are mapped as US-specific considerations where they add or differ from ISO 13485. - Focus on whether the document (a) defines the required process/structure and (b) shows it is practicable/implemented. Do not assess actual design outputs or project-specific records—only the procedural framework. - For combination products: assess whether the SOP addresses 21 CFR 4 CGMP applicability, drug-device interface considerations, and coordination with drug development. How to review each clause invocation: 1) Understand what type of document this is and how it fits the Quality Management System. 2) Focus ONLY on the requested clause; search the entire document for relevant evidence. 3) Presence vs adequacy: assess basic alignment with the clause. 4) Cross-references: if the document points to another controlled document (e.g., Risk Management Procedure, Usability Engineering Procedure), consider that evidence. Decision logic: - PASS: Requirement is clearly addressed; process/structure is described with sufficient detail for implementation. - FLAGGED: Evidence exists but is incomplete/ambiguous or needs human confirmation. - FAIL: Core expectations are missing/contradicted. - NOT_APPLICABLE: Clause truly does not apply to the document provided. Regulatory alignment notes: - ISO 13485:2016 Section 7.3 is harmonized with FDA QMSR (effective Feb 2026) - 21 CFR 820.30 legacy requirements noted where they provide additional specificity - EU MDR requires ISO 13485 compliance; technical documentation per Annex II - For combination products, 21 CFR 4 determines which CGMP requirements apply
Example Prompt:
You are an expert medical device risk-management assessor with deep knowledge of ISO 14971:2019 and ISO/TR 24971. Review ONE DOCUMENT AT A TIME and judge whether it addresses specific requirements from ISO 14971:2019 clauses 4-10. Context and assumptions: - Treat the document as a top-level risk-management artifact (procedure, RMP, RMR, etc.). It may reference other SOPs, work instructions, or records; clear cross-references are acceptable evidence that such systems/records exist. - Focus on whether the document (a) defines the required process/structure and (b) shows it is practicable/implemented. Do not score clauses 1-3 or annexes as standalone requirements. How to review each clause invocation: 1) Understand what type of document this is and how it fits the risk-management system. 2) Focus ONLY on the requested clause; search the entire document (headings, lists, tables, appendices, images/OCR) for relevant evidence of the process and expected records. 3) Presence vs adequacy: assess basic alignment with the clause. Minor ambiguity or “could be better” solutions can still PASS; treat those as opportunities for improvement (OFIs). 4) Cross-references: if the document points to another controlled SOP or record, consider that evidence that the process/record exists; do not invent details that are not written. Decision logic: - PASS: Requirement is clearly addressed; process/structure is described. - FLAGGED: Evidence exists but is incomplete/ambiguous or needs human confirmation. - FAIL: Core expectations are missing/contradicted. - NOT_APPLICABLE: Clause truly does not apply to the document provided.

## Req Prompt_Detailed (AI)
Section | Purpose | Writing Tips
Objective | What the clause fundamentally requires | Start with "Determine whether..." - be specific
Core evidence | Minimum threshold for PASS | These are the "must haves" - if ALL present, it PASSES
Strong PASS | Concrete patterns to recognize | Use specific document artifacts
When to FLAG | Uncertainty conditions | Evidence exists but incomplete, ambiguous, or conflicting
When to FAIL | Missing/contradicted | Core expectations are absent
Typical OFIs | Improvement opportunities | Things that could be better but don't prevent PASS
ID/Clause Number | ID/Clause Title | Objective Description | Objective Bullets Add as many as appropriate | Minimum Subset to Pass Bullets Add as many as appropriate | Examples - Strong Pass Bullets Add as many as appropriate | Conditions - Flag Bullets Add as many as appropriate | Conditions - Fail Bullets Add as many as appropriate | Conditions - OFI Bullets Add as many as appropriate
ISO7.3.1-01 | General Requirements - Design and Development Procedures | Determine whether the organization has established documented procedures for design and development. | - Documented procedures for design and development established - Procedures address all aspects of Section 7.3 - Procedures maintained and controlled | - Design and development procedure(s) exist and are documented - Procedures cover planning through transfer - Document control applied to procedures - Scope addresses medical device design activities | - Comprehensive procedure covering all 7.3 subclauses - Clear scope and applicability statements - Defined terms and abbreviations - Revision history and approval signatures | - Procedures exist but incomplete coverage of 7.3 - Some elements addressed in separate unlinked documents - Scope unclear or overly broad - Document control status uncertain | - No documented design and development procedures - Procedures exist but not controlled documents - Major gaps in 7.3 coverage - Procedures not appropriate for device risk | - Process flowchart for visual overview - Linkage to other QMS procedures - Training requirements for procedure users - Procedure effectiveness metrics
ISO7.3.2-01 | Design Planning - Design and Development Planning Requirements | Determine whether the organization plans and controls design and development of the medical device. | - Design and development planning required - Planning includes control of design activities - Planning appropriate to nature of device - Planning documented and maintained | - Design planning requirements established - Roles and responsibilities for design control activities defined - Planning scaled to device complexity/risk - Plan documentation requirements specified | - Design plan template provided - Scalable planning approach for different device classes - Planning triggers and timing defined - Plan as living document with update requirements | - Planning required but format/content unclear - Roles and responsibilities vague - Planning not appropriate for device complexity/risk - Plan documentation location not specified | - No design planning requirements - Planning informal or ad hoc - No control over design activities - Plans not documented | - Agile/iterative planning options - Planning templates for common device types - Planning metrics and KPIs
ISO7.3.2-02 | Design Planning - Design and Development Stages | Determine whether design planning includes defined design and development stages. | - Design stages/phases defined - Stage boundaries clear - Stage-appropriate activities identified - Progression criteria between stages | - Design stages explicitly defined in procedure - Clear description of each stage purpose - Activities mapped to stages - Stage transition criteria (phase gates, design reviews) | - Named stages (e.g., Concept, Development, Verification, Validation, Transfer) - Entry/exit criteria for each stage - Deliverables defined per stage - Decision authority for stage progression | - Stages mentioned without clear definition - Activities not mapped to stages - Transition criteria vague - Stage flexibility/tailoring not addressed | - No design stages defined - Linear process without checkpoints - No stage-based control - Activities not sequenced | - Stage gate review checklists - Concurrent engineering considerations - Fast-track options for low-risk changes - Stage duration guidelines
ISO7.3.2-03 | Design Planning - Review, Verification, and Validation Planning | Determine whether planning includes review, verification, and validation appropriate at each stage. | - Reviews planned at appropriate stages - Verification activities planned per stage - Validation activities planned per stage - V&V distinct and appropriately timed | - Review timing linked to stages - Verification activities mapped to design phases - Validation timing defined (late stage, production units) - Clear distinction between V&V | - V&V matrix by design stage - Review types defined (technical, management, phase gate) - Resource planning for V&V activities - V&V deliverables per stage | - V&V mentioned but not mapped to stages - Review timing unclear - V&V distinction not clear - Resource needs not addressed | - No V&V planning requirements - V&V treated as end-of-project only - No review planning - V&V conflated | - Risk-based V&V planning and traceability - Sample size planning guidance - V&V protocol templates - Supplier V&V coordination
ISO7.3.2-04 | Design Planning - Responsibilities and Authorities | Determine whether planning includes responsibilities and authorities for design and development. | - Design responsibilities defined - Authorities for decisions assigned - Roles documented in plans - Competency requirements addressed | - Responsibility assignment required in plans - Decision authority defined - Role documentation requirements - Competency linkage to training | - RACI matrix or equivalent required - Authority levels for design decisions - Cross-functional role definitions - Competency matrix for design roles | - Responsibilities mentioned without method - Authority levels unclear - Role definitions vague - Competency not linked | - No responsibility requirements - Authorities undefined - Roles not documented - Competency not addressed | - Delegation of authority process - External resource responsibilities - Responsibility transfer at phase gates
ISO7.3.2-05 | Design Planning - Internal and External Interfaces | Determine whether planning addresses management of interfaces between different groups involved in design. | - Internal interfaces identified - External interfaces identified - Interface management defined - Communication across interfaces ensured | - Interface identification required - Internal groups (R&D, QA, RA, Mfg) addressed - External parties (suppliers, consultants) addressed - Communication mechanisms defined | - Interface matrix required in plans - Meeting/coordination requirements - Information exchange protocols - Escalation paths for interface issues | - Interfaces mentioned without specifics - Internal only or external only addressed - Communication informal - Escalation not defined | - No interface management - Design treated as isolated activity - No communication requirements - Coordination not addressed | - Drug development team interface for combination products - Notified body/FDA interface - Clinical site interface - Contract manufacturer interface
ISO7.3.2-06 | Design Planning - Traceability Methods | Determine whether planning includes methods to ensure traceability of design outputs to inputs. | - Traceability methodology required - Output-to-input linkage ensured - Traceability tools/methods defined - Traceability maintained throughout design | - Traceability requirement explicit - Method for maintaining traceability defined - Bidirectional traceability addressed - Traceability review/verification | - Traceability matrix required - Tool requirements (spreadsheet, software) - Traceability verification at reviews - Gap analysis requirements | - Traceability mentioned without method - Tools not specified - Verification not required - Bidirectional not addressed | - No traceability requirements - Inputs/outputs disconnected - No method defined - No verification | - Electronic traceability tools - Traceability to risk controls - Traceability to test cases - Traceability audit guidance
ISO7.3.2-07 | Design Planning - Resource Planning | Determine whether planning includes resources needed, including necessary competence of personnel. | - Resource planning required - Personnel competence addressed - Equipment/tools considered - Resource availability assessed | - Resource requirements in plans - Competence requirements defined - Resource allocation process - Competence verification | - Resource plan template - Competence matrix reference - Gap analysis and mitigation - External resource qualification | - Resources mentioned without process - Competence vague - Allocation informal - Gaps not addressed | - No resource planning - Competence not addressed - Resources assumed available - No allocation process | - Resource loading/leveling - Training plan integration - Contingency planning - Budget linkage
ISO7.3.2-08 | Design Planning - Plan Updates | Determine whether the procedure requires design plans to be updated as design progresses. | - Plan update requirements defined - Update triggers identified - Approval for updates required - Plan history maintained | - Plan updates required as design evolves - Triggers for updates defined - Update approval process - Version control required | - Specific update triggers (scope change, schedule, findings) - Update review and approval workflow - Change documentation requirements - Plan baseline management | - Updates mentioned without triggers - Approval process unclear - Version control vague - Baseline not addressed | - No plan update requirements - Plans treated as static - No approval for changes - No version control | - Plan update metrics - Automated update reminders - Stakeholder notification - Lessons learned integration
ISO7.3.2-US01 | Design Planning - US: Design History File Planning (820.30) | Determine whether planning addresses Design History File requirements per 21 CFR 820.30. | - DHF requirements addressed in planning - DHF structure/index planned - DHF ownership assigned - DHF completeness criteria defined | - DHF referenced in planning requirements - DHF content planning addressed - DHF responsibility assigned - Relationship to design plan defined | - DHF index template - DHF building throughout project - DHF review at phase gates - DHF completeness checklist | - DHF mentioned but planning unclear - Content not specified - Timing of DHF activities vague - Ownership unclear | - No DHF planning - DHF treated as afterthought - No ownership - No content requirements | - Electronic DHF system - DHF audit readiness - DHF for device families - Legacy DHF guidance
ISO7.3.2-CP01 | Design Planning - Combination Product: CGMP Determination (21 CFR 4) | Determine whether planning addresses 21 CFR 4 CGMP applicability for combination products. | - 21 CFR 4 applicability assessed - Subpart A vs B determination documented - Constituent parts identified - CGMP responsibilities assigned | - CGMP determination required for combination products - Constituent part identification process - Rationale documentation required - Cross-functional input for determination | - Decision tree for CGMP determination - Rationale template - Regulatory strategy linkage - Periodic reassessment triggers | - Combination products acknowledged but 21 CFR 4 vague - Determination process unclear - Documentation requirements missing - Coordination informal | - No combination product consideration - 21 CFR 4 not referenced - No constituent identification - CGMP assumed without assessment | - FDA pre-submission strategy - Examples of Subpart A/B scenarios - Drug development coordination SOP - CGMP election change management
ISO7.3.2-RM01 | Design Planning - Risk Management Planning Integration (ISO 14971) | Determine whether design planning integrates risk management planning per ISO 14971. | - Risk management plan required - Risk activities mapped to design stages - Risk acceptability criteria referenced - Risk file requirements established | - RMP required for design projects - Risk activities linked to design phases - Acceptability criteria defined or referenced - Risk file maintenance requirements | - RMP template or content requirements - Timing of risk activities vs design milestones - ALARP methodology referenced - Risk review at design reviews | - Risk mentioned but planning integration unclear - Activities not mapped to stages - Acceptability criteria not specified - Risk file location undefined | - No risk management planning - Risk analysis disconnected from design - No integration with design phases - Risk file not required | - Use-related risk analysis integration - Production risk analysis timing - Post-production monitoring linkage - Supplier risk coordination
ISO7.3.2-UE01 | Design Planning - Usability Engineering Planning (IEC 62366-1) | Determine whether design planning integrates usability engineering per IEC 62366-1. | - Use specification requirements - UE process integrated with design - Formative/summative study milestones - User population documentation | - Use specification required - UE activities mapped to design phases - Formative evaluation in early phases - Summative validation in late phases | - Use specification template - HF study milestones in design plan - User population characterization - Critical task identification process | - UE mentioned but planning unclear - Formative/summative not distinguished - Use specification content undefined - User population vague | - No UE planning requirements - Usability ad hoc - No use specification - User needs not in planning | - Known use problems analysis - Training effectiveness planning - Labeling development integration - HF specialist resource planning
ISO7.3.3-01 | Design Input - Design Input Determination | Determine whether inputs relating to product requirements are determined and records maintained. | - Product requirement inputs determined - Input sources identified - Input determination process defined - Records of inputs maintained | - Input determination process established - Multiple input sources addressed - Documentation requirements defined - Record retention specified | - Input sources checklist (user, regulatory, standards, risk) - Input gathering methodology - Input documentation template - Traceability initiation at input stage | - Inputs required but sources vague - Determination process unclear - Documentation informal - Record requirements incomplete | - No input determination process - Sources not identified - Inputs undocumented - No records maintained | - Stakeholder analysis for inputs - Competitive analysis as input - Clinical advisory input - Post-market feedback integration
ISO7.3.3-02 | Design Input - Functional and Performance Requirements | Determine whether design inputs include functional and performance requirements. | - Functional requirements as inputs - Performance requirements as inputs - Requirements measurable/verifiable - Requirements complete for intended use | - Functional requirements explicitly required - Performance specifications required - Measurability criteria addressed - Completeness review required | - Functional requirement format guidance - Performance specification template - Quantitative vs qualitative guidance - Acceptance criteria at input stage | - Requirements mentioned without categorization - Measurability not addressed - Completeness criteria vague - Format not specified | - No functional requirements - No performance specifications - Requirements not verifiable - Input completeness not assessed | - Essential performance per IEC 60601 - Worst-case performance conditions - Degradation requirements - Environmental performance
ISO7.3.3-03 | Design Input - Safety Requirements | Determine whether design inputs include safety requirements. | - Safety requirements as inputs - Safety derived from risk management - Safety standards addressed - Patient and user safety covered | - Safety requirements explicitly required - Risk-derived safety requirements - Applicable safety standards identified - User and patient safety distinguished | - Safety requirement derivation process - Linkage to preliminary hazard analysis - Safety standard checklist - Safety requirement format | - Safety mentioned without process - Risk linkage unclear - Standards not specified - Patient/user not distinguished | - No safety requirements - Safety disconnected from risk - Standards not addressed - Safety assumed | - Use-related safety requirements - Cybersecurity requirements - Biocompatibility requirements - EMC/electrical safety
ISO7.3.3-04 | Design Input - Applicable Regulatory Requirements | Determine whether design inputs include applicable regulatory requirements. | - Regulatory requirements as inputs - Multiple markets addressed - Standards applicability assessed - Regulatory pathway impact | - Regulatory requirements required as inputs - Market-specific requirements identified - Standards matrix required - Regulatory traceability | - Regulatory input checklist by market - Standards applicability matrix - Classification-specific requirements - Predicate/essential equivalence | - Regulatory inputs mentioned without specifics - Single market focus only - Standards informal - Traceability unclear | - No regulatory input requirements - Regulations assumed, not documented - Standards not identified - No regulatory traceability | - Emerging regulation monitoring - Guidance document review - Special controls identification - International harmonization
ISO7.3.3-05 | Design Input - Risk Management Outputs as Inputs | Determine whether design inputs include applicable outputs of risk management. | - Risk outputs as design inputs - Hazard mitigation requirements - Risk control specifications - Risk-input traceability | - Risk management outputs required as inputs - Hazard analysis feeds design requirements - Risk controls specified as inputs - Traceability to risk analysis | - Preliminary hazard analysis timing - Risk-derived requirement format - Safety requirement derivation - Bidirectional traceability | - Risk inputs mentioned without process - Timing unclear - Format not specified - Traceability vague | - No risk-derived inputs - Risk analysis disconnected - Safety requirements ad hoc - No traceability | - Use-related risk inputs - Production risk inputs - Post-market feedback integration - Supplier risk inputs
ISO7.3.3-06 | Design Input - Information from Previous Designs | Determine whether design inputs include information derived from previous similar designs, as appropriate. | - Previous design information considered - Predicate/similar device analysis - Lessons learned incorporated - Known issues addressed | - Previous design review required - Predicate analysis for regulated devices - Lessons learned process - Known problem resolution | - Predicate comparison template - Lessons learned database - Known use error analysis - Field performance data review | - Previous designs mentioned without process - Predicate analysis informal - Lessons learned ad hoc - Known issues not tracked | - No previous design consideration - Predicate not analyzed - Lessons not incorporated - Known issues ignored | - Competitor device analysis - Literature review process - MAUDE/vigilance database review - Patent landscape review
ISO7.3.3-07 | Design Input - Essential Requirements for Design | Determine whether design inputs include other requirements essential for design and development. | - Essential requirements identified - Manufacturing requirements as inputs - Packaging/labeling requirements - Service/maintenance requirements | - Comprehensive input categories - Manufacturing constraints addressed - Packaging/labeling inputs - Lifecycle considerations | - Input category checklist - DFM/DFA requirements - Sterilization requirements - Shelf-life requirements | - Essential requirements vague - Manufacturing informal - Packaging/labeling late - Service not addressed | - Essential requirements missing - Manufacturing disconnected - No packaging/labeling inputs - Lifecycle ignored | - Environmental requirements - Cost targets as inputs - Supply chain requirements - Sustainability requirements
ISO7.3.3-08 | Design Input - Input Review and Adequacy | Determine whether design inputs are reviewed for adequacy and approved. | - Input review required - Adequacy assessment criteria - Incomplete/ambiguous resolution - Approval before use | - Input review process defined - Adequacy criteria established - Resolution process for deficiencies - Approval workflow defined | - Input review checklist - Completeness criteria - Conflict resolution process - Cross-functional review requirement | - Review required but criteria vague - Adequacy assessment informal - Resolution process unclear - Approval authority undefined | - No input review - Adequacy not assessed - Deficiencies not resolved - No approval required | - Input freeze/baseline process - Change management for inputs - Stakeholder sign-off - Input metrics
ISO7.3.3-CP01 | Design Input - Combination Product: Drug Product Requirements as Inputs | Determine whether design inputs include drug product specifications for combination products. | - Drug formulation requirements - Container closure requirements - Stability/compatibility requirements - Dose delivery specifications | - Drug product specs as device inputs - Container closure addressed - Compatibility/stability required - Dose accuracy from drug specs | - Drug-device interface spec template - CMC coordination process - Extractables/leachables consideration - Drug labeling requirements | - Drug requirements mentioned without detail - Container closure vague - Compatibility informal - CMC coordination unclear | - No drug product inputs - Device designed without drug - Container closure missing - Interface not specified | - Shelf-life harmonization - Temperature/storage alignment - Reconstitution requirements - Multi-formulation compatibility
ISO7.3.3-UE01 | Design Input - User Interface Requirements as Inputs (IEC 62366-1) | Determine whether design inputs include user interface and use-related requirements. | - User interface characteristics - Use-related hazards identified - Use scenarios documented - Critical task requirements | - UI requirements as inputs - Use-related hazard identification - Use scenario documentation - Critical task identification | - Task analysis methodology - Use error analysis from predicates - User population-specific requirements - Use environment requirements | - UI mentioned without process - Use scenarios vague - Critical tasks undefined - Hazard identification informal | - No UI input requirements - Use scenarios missing - Task analysis absent - Use-related hazards ignored | - Competitive usability benchmarking - User feedback integration - Simulated use environment specs - Training requirement inputs
ISO7.3.4-01 | Design Output - Output Form for Verification | Determine whether design outputs are in a form suitable for verification against inputs. | - Outputs in verifiable form - Outputs traceable to inputs - Measurable/testable outputs - Clear acceptance criteria | - Output format requirements defined - Verifiability required - Traceability to inputs required - Acceptance criteria before verification | - Output specification template - Measurability criteria - Traceability matrix requirements - Acceptance criteria format | - Outputs required but format unclear - Verifiability not explicit - Traceability mentioned without method - Acceptance criteria timing vague | - No output format requirements - Outputs not verifiable - No input linkage - No acceptance criteria | - Statistical acceptance criteria - Pass/fail vs variable data - Worst-case condition specification - Output categorization
ISO7.3.4-02 | Design Output - Output Approval Before Release | Determine whether design outputs are approved prior to release. | - Output approval required - Approval before release/use - Approval authority defined - Approval documented | - Approval requirement explicit - Timing before release defined - Authority designated - Documentation required | - Approval workflow defined - Multi-level approval as needed - Electronic signature compliance - Release criteria checklist | - Approval mentioned without process - Timing unclear - Authority vague - Documentation informal | - No approval requirements - Outputs released without control - Authority undefined - No documentation | - Conditional release process - Partial approval criteria - Approval delegation - Approval metrics
ISO7.3.4-03 | Design Output - Purchasing, Production, and Service Information | Determine whether design outputs provide appropriate information for purchasing, production, and service provision. | - Purchasing information as outputs - Production specifications - Service/maintenance information - DMR content addressed | - Manufacturing specifications required - Purchasing specifications required - Service information as outputs - DMR linkage defined | - DMR content checklist - Drawing/specification requirements - Process specification format - Supplier specification requirements | - Manufacturing outputs vague - Purchasing specs not explicit - Service information missing - DMR relationship unclear | - No manufacturing outputs - Design-production disconnect - No purchasing specifications - No service information | - Process validation linkage - Supplier qualification coordination - Assembly/test instructions - Packaging specifications
ISO7.3.4-04 | Design Output - Product Acceptance Criteria | Determine whether design outputs contain or reference product acceptance criteria. | - Acceptance criteria in outputs - Criteria for product release - Inspection/test criteria - Criteria traceable to inputs | - Acceptance criteria required in outputs - Product release criteria defined - Inspection/test specifications - Input traceability | - Acceptance criteria format - Statistical basis where appropriate - In-process and final criteria - Sampling plan references | - Criteria mentioned without format - Release criteria vague - Test specs informal - Traceability unclear | - No acceptance criteria - Release criteria undefined - No test specifications - Criteria disconnected | - AQL/sampling guidance - Attribute vs variable criteria - Supplier acceptance criteria - Stability acceptance criteria
ISO7.3.4-05 | Design Output - Essential Characteristics for Safe and Proper Use | Determine whether design outputs specify characteristics essential for safe and proper use. | - Essential characteristics identified - Safety-critical outputs defined - Proper use characteristics - Risk control outputs | - Essential characteristics required - Safety linkage explicit - Proper use specifications - Risk control documentation | - Essential characteristics list - Risk-based prioritization - Critical-to-quality identification - Regulatory linkage (substantial equivalence) | - Essential mentioned without criteria - Safety linkage vague - Prioritization unclear - Risk connection informal | - No essential characteristics - Safety outputs undefined - No prioritization - Risk controls not outputs | - Essential performance testing - Process capability for essentials - Change control for essentials - Post-market monitoring linkage
ISO7.3.4-06 | Design Output - Risk Control Measures as Outputs (ISO 14971) | Determine whether design outputs include risk control measures. | - Risk controls documented as outputs - Control specifications - Residual risk documentation - Risk-output traceability | - Risk controls in output documentation - Control specifications defined - Residual risk documented - Traceability to risk analysis | - Risk control specification format - Residual risk acceptance documentation - Verification protocol linkage - Risk management report integration | - Risk controls mentioned without format - Specifications vague - Residual risk informal - Traceability unclear | - No risk control outputs - Controls not specified - Residual risk undocumented - No traceability | - Use-related control outputs - Information for safety outputs - Training as control output - Post-market control effectiveness
ISO7.3.4-07 | Design Output - User Interface Specification as Output (IEC 62366-1) | Determine whether design outputs include user interface specification. | - UI specification as output - Safety-relevant UI elements - Labeling/IFU as outputs - Training materials | - UI specification required - Safety elements identified - Labeling as design output - Training requirements | - UI specification template - IFU content requirements - Training effectiveness criteria - Usability test acceptance criteria | - UI mentioned without detail - Labeling not explicit output - Training not addressed - Safety elements vague | - No UI specification - Labeling uncontrolled - No training requirements - Use safety not in outputs | - Multilingual requirements - Quick reference guide - Video/multimedia outputs - User feedback mechanism
ISO7.3.5-01 | Design Review - Systematic Reviews at Suitable Stages | Determine whether systematic reviews of design and development are performed at suitable stages. | - Systematic reviews required - Reviews at suitable stages - Review methodology defined - Stage-appropriate scope | - Systematic review requirements - Stage definitions for reviews - Review methodology defined - Scope per stage | - Review types defined (technical, management, gate) - Stage-specific review criteria - Review agenda templates - Phase gate criteria | - Reviews required but stages unclear - Methodology vague - Scope informal - Systematic not defined | - No systematic reviews - Reviews ad hoc - No stage linkage - Methodology undefined | - Review effectiveness metrics - Pre-review package requirements - Review action tracking - Lessons learned capture
ISO7.3.5-02 | Design Review - Evaluation of Design Results | Determine whether reviews evaluate ability of design results to meet requirements. | - Results evaluated against requirements - Input-output conformance assessed - V&V status reviewed - Risk status evaluated | - Requirement conformance evaluation - Design status assessment - V&V progress review - Risk management status | - Review checklist against requirements - Traceability matrix review - V&V summary presentation - Risk report presentation | - Evaluation mentioned without criteria - Conformance assessment vague - V&V status informal - Risk review not explicit | - No evaluation against requirements - Conformance not assessed - V&V not reviewed - Risk not evaluated | - Quantitative progress metrics - Technical debt assessment - Schedule/resource review - Quality metrics review
ISO7.3.5-03 | Design Review - Problem Identification and Actions | Determine whether reviews identify problems and propose necessary actions. | - Problem identification required - Actions proposed - Action tracking - Closure verification | - Problem identification in reviews - Action proposal process - Tracking mechanism - Closure documentation | - Problem categorization (critical, major, minor) - Action ownership and due dates - Escalation criteria - Closure verification process | - Problems mentioned without process - Actions informal - Tracking vague - Closure not verified | - No problem identification - No action tracking - Issues not documented - No closure process | - CAPA linkage for systemic issues - Trend analysis - Risk register update triggers - Design decision documentation
ISO7.3.5-04 | Design Review - Review Participants | Determine whether reviews include representatives of functions concerned with the design stage. | - Function representatives required - Stage-appropriate participation - Specialist inclusion - Independent perspective | - Cross-functional attendance required - Functions defined by stage - Specialist determination - Independent reviewer (per 820.30) | - Attendee matrix by review type - Quorum requirements - Specialist identification criteria - Independence definition | - Participation required but vague - Specialists not addressed - Independence unclear - Attendance not documented | - No participation requirements - Single function acceptable - No specialist requirements - No independence | - Drug development for combination products - HF specialist requirements - Clinical input requirements - Supplier participation
ISO7.3.5-05 | Design Review - Review Records | Determine whether records of review results and necessary actions are maintained. | - Review results recorded - Actions documented - Records maintained - Traceability to design file | - Results documentation required - Action documentation required - Record retention defined - Design file linkage | - Review minutes template - Action item format - Attendee/signature records - DHF/design file requirements | - Documentation required but format unclear - Actions informal - Retention vague - File linkage not specified | - No record requirements - Results not documented - Actions not recorded - No design file | - Electronic review records - Action closure tracking - Review effectiveness metrics - Searchable records
ISO7.3.5-06 | Design Review - Risk Management Review Integration (ISO 14971) | Determine whether design reviews include risk management status review. | - Risk status at reviews - Risk file adequacy - Control implementation - Residual risk evaluation | - Risk review as agenda item - Risk file assessment - Control status review - Residual risk discussion | - Risk review checklist - Risk management report presentation - Open risk item tracking - Risk acceptance documentation | - Risk mentioned without specifics - File review not explicit - Control status vague - Residual risk not discussed | - No risk review - Risk separate from design reviews - Status not assessed - No risk acceptance | - Use-related risk review - Production risk review - Post-market risk integration - Cumulative risk assessment
ISO7.3.5-07 | Design Review - Human Factors Review Integration (IEC 62366-1) | Determine whether design reviews include human factors/usability evaluation status. | - HF status at reviews - Formative study results - Use error analysis - UI adequacy assessment | - HF review as agenda item - Formative results presented - Use errors discussed - UI adequacy evaluated | - HF specialist presentation - Use error trending - Summative readiness assessment - Critical task review | - HF mentioned without requirements - Formative not reviewed - Use errors not tracked - UI criteria vague | - No HF review - Usability separate - Use errors ignored - UI not assessed | - User feedback review - Training effectiveness - Labeling adequacy - Competitive comparison
ISO7.3.6-01 | Design Verification - Verification per Planned Arrangements | Determine whether design verification is performed in accordance with planned arrangements. | - Verification per plan - Planned arrangements defined - Verification methods appropriate - Timing per design stage | - Verification per planning requirements - Methods defined in plan - Stage-appropriate verification - Protocol requirements | - Verification matrix per plan - Protocol template - Method selection criteria - Statistical rationale requirements | - Verification required but planning link unclear - Methods vague - Timing not specified - Protocol format undefined | - No verification planning link - Methods ad hoc - Uncontrolled verification - No protocols | - Verification vs validation distinction - Sample size guidance - Worst-case selection - Accelerated testing guidance
ISO7.3.6-02 | Design Verification - Output Conformance to Inputs | Determine whether verification ensures design outputs meet design input requirements. | - Output-input conformance verified - Traceability demonstrated - All inputs verified - Gap identification | - Conformance verification required - Traceability matrix verification - Input coverage assessment - Gap analysis process | - Verification report showing traceability - 100% input coverage requirement - Gap closure documentation - Summary verification matrix | - Conformance mentioned without method - Traceability informal - Coverage not assessed - Gaps not tracked | - No conformance verification - Inputs-outputs disconnected - Coverage unknown - Gaps ignored | - Partial verification criteria - Negative testing - Boundary condition testing - Regression verification
ISO7.3.6-03 | Design Verification - Verification Records | Determine whether records of verification results and necessary actions are maintained. | - Verification results recorded - Actions documented - Records complete - Design file integration | - Result documentation required - Action documentation - Record completeness criteria - Design file requirements | - Verification report template - Protocol-report linkage - Deviation handling - DHF integration | - Documentation required but incomplete - Actions informal - Completeness vague - File integration unclear | - No verification records - Results undocumented - Actions not tracked - No design file | - Electronic data integrity - Raw data retention - OOS handling - Summary reports
ISO7.3.6-04 | Design Verification - Risk Control Verification (ISO 14971) | Determine whether verification includes risk control implementation and effectiveness. | - Control implementation verified - Effectiveness verified - Verification methods for controls - Documentation requirements | - Implementation verification required - Effectiveness verification required - Methods defined - Documentation specified | - Control verification matrix - Effectiveness criteria - Risk file linkage - Residual risk verification | - Verification mentioned without method - Effectiveness vague - Documentation unclear - Residual risk not addressed | - No control verification - Controls assumed effective - No documentation - Risk-verification disconnected | - Use-related control verification - Information for safety verification - Training verification - Post-market monitoring
ISO7.3.6-05 | Design Verification - Formative Usability Evaluation (IEC 62366-1) | Determine whether formative usability evaluation is included in design verification. | - Formative evaluation required - Methods defined - Iterative improvement - Design feedback mechanism | - Formative evaluation requirements - Method options (expert review, user testing) - Iteration process - Design change triggers | - Formative protocol guidance - User recruitment criteria - Finding categorization - Design improvement tracking | - Formative mentioned without requirements - Methods vague - Iteration unclear - Feedback informal | - No formative evaluation - Usability only at end - No iteration - Findings not used | - Remote vs in-person - Prototype fidelity - Participant numbers - Comparative evaluation
ISO7.3.6-CP01 | Design Verification - Combination Product: Verification with Drug Product | Determine whether verification includes testing with drug product for combination products. | - Drug product in verification - Functional performance with drug - Dose accuracy verification - Container closure integrity | - Drug product use required - Functional testing with drug - Dose accuracy testing - Container closure testing | - Surrogate/bridging rationale - Drug lot selection criteria - Dose accuracy acceptance criteria - Stability condition testing | - Drug use mentioned without requirements - Surrogate without rationale - Criteria unclear - Container closure vague | - No drug product verification - Device tested alone - Dose accuracy missing - Container closure absent | - Drug lot variability - Aged drug testing - Temperature excursion - Multi-formulation compatibility
ISO7.3.7-01 | Design Validation - Validation per Planned Arrangements | Determine whether design validation is performed in accordance with planned arrangements. | - Validation per plan - Planned arrangements defined - Validation timing appropriate - User needs focus | - Validation per planning - Methods in plan - Late-stage timing - User needs linkage | - Validation matrix per plan - Protocol template - Timing criteria - User needs traceability | - Validation required but planning vague - Methods unclear - Timing not specified - User needs link informal | - No validation planning link - Methods ad hoc - Timing inappropriate - User needs ignored | - Clinical study integration - Real-world evidence - User satisfaction - Post-market validation
ISO7.3.7-02 | Design Validation - Conformance to User Needs and Intended Uses | Determine whether validation ensures product conforms to defined user needs and intended uses. | - User needs validated - Intended use demonstrated - Actual/simulated conditions - User population tested | - User needs traceability - Intended use validation - Use condition requirements - Representative users | - User needs mapping in protocol - Intended use acceptance criteria - Use environment specification - User group definitions | - User needs mentioned without method - Intended use vague - Conditions unclear - Users not defined | - No user needs validation - Intended use not demonstrated - Conditions inappropriate - Users not representative | - Patient vs HCP validation - Home vs clinical use - Training effectiveness - Long-term use validation
ISO7.3.7-03 | Design Validation - Validation Prior to Release | Determine whether validation is completed prior to product release for use or delivery. | - Validation before release - Completion criteria - Release decision linkage - Timing requirements | - Validation timing before release - Completion definition - Release criteria include validation - Timing enforced | - Validation as release gate - Completion checklist - Release decision documentation - Partial release criteria | - Timing mentioned without enforcement - Completion vague - Release link unclear - Exceptions undefined | - No timing requirement - Validation after release acceptable - Completion undefined - Release disconnected | - Conditional release - Post-market commitments - Validation completion metrics - Regulatory submission timing
ISO7.3.7-04 | Design Validation - Validation Records | Determine whether records of validation results and necessary actions are maintained. | - Validation results recorded - Actions documented - Records complete - Design file integration | - Result documentation required - Action documentation - Completeness criteria - Design file requirements | - Validation report template - Protocol-report linkage - Statistical analysis requirements - Regulatory submission format | - Documentation incomplete - Actions informal - Completeness vague - File integration unclear | - No validation records - Results undocumented - Actions not tracked - No design file | - Electronic records - Raw data retention - Summary reports - Trend analysis
ISO7.3.7-05 | Design Validation - Summative Usability Validation (IEC 62366-1) | Determine whether summative human factors validation study is required. | - Summative study required - Final design/labeling/training - Representative users - Use safety evidence | - Summative requirements - Final design requirement - Representative user requirement - Safety evidence requirements | - Summative protocol template - User group definitions - Sample size (15+ per group) - Acceptance criteria (no critical task failures) | - Summative mentioned without requirements - Final design unclear - User representation vague - Criteria undefined | - No summative requirements - Design not finalized - Users not representative - No use safety validation | - Simulated use environment - Interview methodology - Use error RCA - Knowledge task assessment
ISO7.3.7-06 | Design Validation - Overall Residual Risk Evaluation (ISO 14971) | Determine whether overall residual risk evaluation is performed before product release. | - Overall risk evaluated - Risk-benefit determination - Before release timing - Acceptability documented | - Overall evaluation required - Risk-benefit if needed - Timing specified - Acceptability documented | - Risk management report requirements - Risk-benefit methodology - Stakeholder involvement - Residual risk disclosure | - Evaluation mentioned without process - Risk-benefit vague - Timing unclear - Documentation incomplete | - No overall evaluation - Individual risks only - No risk-benefit - Release without assessment | - Clinical input - Patient perspective - Comparative assessment - Post-market monitoring plan
ISO7.3.7-CP01 | Design Validation - Combination Product: Validation with Drug Product | Determine whether validation uses commercial-equivalent drug product for combination products. | - Commercial-equivalent drug - Clinical performance validation - Combination use validated - Patient population | - Drug product requirements - Clinical study linkage - Combination use requirements - Patient population requirements | - Drug lot selection criteria - Clinical protocol integration - Real-world conditions - Patient subgroups | - Drug use mentioned without requirements - Clinical linkage vague - Combination use unclear - Population undefined | - No drug product requirements - Clinical disconnected - Device-only validation - No patient studies | - Drug stability at use - Self-administration - HCP administration - Home vs clinical
ISO7.3.8-01 | Design Transfer - Transfer Procedures | Determine whether procedures for transfer of design outputs to manufacturing are documented. | - Transfer procedures documented - Output to manufacturing translation - Transfer completeness - Transfer verification | - Transfer procedure requirements - Translation process defined - Completeness criteria - Verification of transfer | - Transfer checklist - DMR completeness verification - Process validation linkage - First article inspection | - Transfer mentioned without procedure - Translation vague - Completeness undefined - Verification not specified | - No transfer procedures - Design-manufacturing gap - No completeness criteria - No verification | - Supplier transfer - Training documentation - Equipment qualification - Scale-up process
ISO7.3.8-02 | Design Transfer - Manufacturing Suitability Verification | Determine whether design outputs are verified as suitable for manufacturing prior to release. | - Suitability verified - Manufacturing capability assessed - Production requirements met - Pre-production verification | - Suitability verification required - Capability assessment - Requirement conformance - Pre-production activities | - Capability study requirements - Cpk/Ppk criteria - Equipment qualification - Operator training | - Suitability mentioned without criteria - Capability informal - Requirements vague - Pre-production undefined | - No suitability verification - Capability not assessed - Requirements not verified - Transfer uncontrolled | - SPC implementation - Continuous improvement - Supplier capability - Multi-site capability
ISO7.3.8-03 | Design Transfer - Production Capability for Product Requirements | Determine whether production capability can meet product requirements. | - Capability demonstrated - Product requirements met - Process validation - Ongoing capability | - Capability demonstration required - Requirement linkage - Process validation requirements - Capability monitoring | - Process validation protocol - Capability indices - Control plan - Monitoring requirements | - Capability mentioned without method - Validation linkage vague - Control undefined - Monitoring not specified | - No capability demonstration - Requirements not linked - No process validation - No monitoring | - Continuous verification - Revalidation triggers - Supplier capability - Capacity planning
ISO7.3.8-04 | Design Transfer - Labeling and Training Transfer (IEC 62366-1) | Determine whether finalized labeling and training materials are released at transfer. | - Labeling finalized - Training materials released - IFU controlled - Materials validated | - Labeling release required - Training requirements - IFU in DMR - Validation required | - Labeling approval process - Training effectiveness validation - Multi-language requirements - Packaging requirements | - Labeling timing unclear - Training not addressed - IFU control vague - Validation incomplete | - No labeling requirements - Training absent - IFU uncontrolled - No validation | - Label printing qualification - Training monitoring - User feedback - Labeling change control
ISO7.3.9-01 | Design Changes - Change Identification and Records | Determine whether design and development changes are identified and records maintained. | - Changes identified - Records maintained - Change documentation - Configuration management | - Change identification process - Record requirements - Documentation format - Configuration control | - Change request form - Change log requirements - Version control - Baseline management | - Identification vague - Records informal - Documentation incomplete - Configuration unclear | - No change identification - No records - Changes undocumented - No configuration control | - Change metrics - Trending analysis - Impact prediction - Change board efficiency
ISO7.3.9-02 | Design Changes - Change Review, Verification, and Validation | Determine whether changes are reviewed, verified, validated as appropriate, and approved before implementation. | - Changes reviewed - Verification as appropriate - Validation as appropriate - Approval before implementation | - Review requirements - V&V determination process - Approval workflow - Implementation control | - Change classification criteria - V&V matrix for changes - Approval authority matrix - Implementation verification | - Review vague - V&V determination unclear - Approval informal - Implementation uncontrolled | - No review requirements - No V&V for changes - No approval - Changes uncontrolled | - Expedited change process - Risk-based classification - Batch/lot impact - Effectiveness verification
ISO7.3.9-03 | Design Changes - Impact on Constituent Parts and Delivered Product | Determine whether change review includes evaluation of effect on constituent parts and product already delivered. | - Constituent part impact - Delivered product impact - Downstream assessment - Field action determination | - Impact assessment required - Constituent evaluation - Delivered product consideration - Field action criteria | - Impact assessment template - Regulatory filing impact - Customer notification criteria - Recall determination | - Impact mentioned without process - Constituent vague - Delivered product informal - Field action unclear | - No impact assessment - Effects not evaluated - Delivered product ignored - Field action not addressed | - Supply chain impact - Clinical impact - Post-market surveillance - Communication plan
ISO7.3.9-04 | Design Changes - Change Records | Determine whether records of change review results and necessary actions are maintained. | - Review results recorded - Actions documented - Records complete - Design file updated | - Result documentation - Action documentation - Record requirements - File update requirements | - Change record template - Action closure tracking - Design file update checklist - Regulatory impact documentation | - Documentation incomplete - Actions informal - Records vague - File update unclear | - No change records - Results undocumented - Actions not tracked - File not updated | - Electronic records - Audit trail - Trend analysis - Effectiveness review
ISO7.3.9-05 | Design Changes - Risk File Update (ISO 14971) | Determine whether risk management file update is assessed for design changes. | - Risk file update assessed - New hazards identified - Control impact evaluated - Residual risk re-evaluated | - Risk file review required - Hazard identification - Control assessment - Residual risk update | - Risk change impact matrix - New hazard process - Control re-verification - RMR update criteria | - Risk mentioned without process - Hazards informal - Controls vague - Residual risk unclear | - No risk file update - Hazards not reassessed - Controls not evaluated - Risk file static | - Cumulative assessment - Risk trending - Production risk - Post-market integration
ISO7.3.9-06 | Design Changes - HF Impact Assessment (IEC 62366-1) | Determine whether human factors impact is assessed for design changes. | - HF impact assessed - UI change evaluation - Revalidation triggers - Use error consideration | - HF assessment required - UI change process - Revalidation criteria - Use error assessment | - HF change checklist - Revalidation decision tree - Formative for significant changes - Critical task impact | - HF mentioned without process - UI changes vague - Revalidation unclear - Use errors informal | - No HF assessment - UI changes uncontrolled - No revalidation process - Use errors ignored | - Training update triggers - Labeling assessment - User feedback - Comparative evaluation
ISO7.3.9-CP01 | Design Changes - Combination Product: Drug-Device Interface Change Assessment | Determine whether change impact on drug product is assessed for combination products. | - Drug impact assessed - Regulatory filing impact - Drug development coordination - CMC assessment | - Drug impact required - Regulatory pathway impact - Coordination requirements - CMC linkage | - Drug-device change matrix - Supplement determination - Stability impact - Change decision tree | - Drug impact vague - Regulatory unclear - Coordination informal - CMC linkage missing | - No drug assessment - Regulatory not addressed - No coordination - CMC disconnected | - FDA/EMA coordination - Comparability triggers - Clinical bridging - PQ impact
ISO7.3.10-01 | Design Files - Design and Development File Requirements | Determine whether a design and development file is maintained for each device type or family. | - File required per device type - File establishment requirements - File maintenance - File completeness | - File requirement per type/family - Establishment timing - Maintenance process - Completeness criteria | - File structure template - Device type/family criteria - File review process - Completeness checklist | - File required but structure vague - Type/family unclear - Maintenance informal - Completeness undefined | - No file requirements - Device-specific files missing - No maintenance - No completeness | - Electronic file system - File archival - Legacy file guidance - Audit checklist
ISO7.3.10-02 | Design Files - Design Conformity Records | Determine whether the file includes records demonstrating conformity to design and development requirements. | - Conformity records included - Design requirements demonstrated - Traceability complete - Records comprehensive | - Conformity records required - Requirement demonstration - Traceability documentation - Record completeness | - Conformity evidence checklist - Traceability matrix in file - Regulatory requirement mapping - Completeness verification | - Records mentioned without specifics - Demonstration vague - Traceability incomplete - Completeness unclear | - No conformity records - Requirements not demonstrated - Traceability absent - Records incomplete | - Conformity summary - Gap analysis documentation - Regulatory mapping - Audit trail
ISO7.3.10-03 | Design Files - Design Change Records in File | Determine whether the file includes records of design and development changes. | - Change records in file - Change history complete - Impact documentation - Approval records | - Change records required - History maintenance - Impact documentation - Approval inclusion | - Change log in file - Complete change history - Impact assessment records - Approval documentation | - Changes mentioned without file requirements - History incomplete - Impact vague - Approvals informal | - No change records in file - History absent - Impact undocumented - Approvals missing | - Change trending - Effectiveness records - Lessons learned - Change metrics
ISO7.3.10-04 | Design Files - Risk Management File Integration (ISO 14971) | Determine whether risk management file is integrated with or referenced by design file. | - Risk file integrated/referenced - Content per ISO 14971 - Traceability complete - File maintained | - Risk file requirements - Integration specified - Content defined - Traceability required | - Risk file content checklist - RMP, analysis, controls, RMR - Complete traceability - Maintenance process | - Risk file mentioned without content - Integration unclear - Content vague - Traceability incomplete | - No risk file requirements - Risk records scattered - Content undefined - No traceability | - Risk file audit checklist - Electronic system - Device families - Post-market updates
ISO7.3.10-05 | Design Files - Usability Engineering File Integration (IEC 62366-1) | Determine whether usability engineering file is integrated with or referenced by design file. | - UE file integrated/referenced - Content per IEC 62366-1 - Traceability complete - File maintained | - UE file requirements - Integration specified - Content defined - Traceability required | - UE file content per 4.1.1 - Use specification included - Formative/summative records - Risk-usability traceability | - UE file mentioned without content - Integration unclear - Content vague - Traceability incomplete | - No UE file requirements - HF records scattered - Content undefined - No traceability | - UE file audit checklist - Electronic system - Device families - International requirements
ISO7.3.10-US01 | Design Files - US: Design History File (21 CFR 820.30) | Determine whether DHF requirements per 21 CFR 820.30(j) are addressed for US market. | - DHF per device type - Plan conformance demonstrated - 820.30 conformance demonstrated - Records complete | - DHF requirement referenced - Plan compliance records - Regulatory compliance - Completeness checklist | - DHF index template - 820.30 mapping - Completeness verification - Audit readiness | - DHF mentioned without specifics - Plan conformance vague - Regulatory mapping incomplete - Completeness undefined | - No DHF requirements - Plan not demonstrated - Regulatory not addressed - Records incomplete | - DHF vs design file alignment - Electronic DHF - Legacy guidance - QMSR transition
ISO7.3.10-CP01 | Design Files - Combination Product: File Requirements | Determine whether design file addresses combination product-specific requirements including drug CMC cross-reference. | - Combination product addressed - Drug CMC cross-reference - CGMP election documented - Interface documentation | - Combination product requirements - CMC reference process - CGMP rationale - Interface documentation | - Combination product checklist - CMC methodology - 21 CFR 4 compliance - Integrated records | - Combination mentioned without specifics - CMC unclear - CGMP vague - Interface informal | - No combination requirements - CMC not referenced - CGMP undocumented - Interface missing | - Co-packaged products - Drug lot traceability - Clinical linkage - Regulatory correspondence

## Requirements Prompt_Simplified
Section | Purpose | Writing Tips
Objective | What the clause fundamentally requires | Start with "Determine whether..." - be specific
Core evidence | Minimum threshold for PASS | These are the "must haves" - if ALL present, it PASSES
Strong PASS | Concrete patterns to recognize | Use specific document artifacts
When to FLAG | Uncertainty conditions | Evidence exists but incomplete, ambiguous, or conflicting
When to FAIL | Missing/contradicted | Core expectations are absent
Typical OFIs | Improvement opportunities | Things that could be better but don't prevent PASS
ID/Clause Number | ID/Clause Title | Objective Description | Objective Bullets Add as many as appropriate | Minimum Subset to Pass Bullets Add as many as appropriate | Examples - Strong Pass Bullets Add as many as appropriate | Conditions - Flag Bullets Add as many as appropriate | Conditions - Fail Bullets Add as many as appropriate | Conditions - OFI Bullets Add as many as appropriate
ISO7.3.1 | General Requirements | Determine whether the organization has established and maintains documented procedures for design and development that address all requirements of ISO 13485:2016 Section 7.3. | - Documented procedures for design and development exist - Procedures address planning, inputs, outputs, review, verification, validation, transfer, changes, and files - Procedures appropriate to device classification and risk level - Procedures maintained under document control - Scope and applicability clearly defined | - Design and development procedure(s) documented and controlled - All 7.3 subclauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled | - Comprehensive procedure with clear structure covering all 7.3 elements - Process flowchart providing visual overview - Defined terms and role definitions - Cross-references to related procedures (Risk Management, Human Factors) - Training requirements identified | - Procedures exist but incomplete coverage of 7.3 subclauses - Some elements in separate documents without clear linkage - Scope unclear or overly broad/narrow - Document control status uncertain | - No documented design and development procedures - Major gaps in 7.3 coverage (e.g., no validation, no design review) - Procedures not controlled documents - Procedures clearly inappropriate for device risk | - Linkage to other QMS processes (CAPA, Supplier, Management) - Examples or worked scenarios - Electronic workflow integration
ISO7.3.2 | Design and Development Planning | Determine whether the organization plans and controls design and development including stages, activities, responsibilities, interfaces, traceability, resources, and plan updates. | - Design and development planning required - Planning includes design control activities - Planning appropriate to nature of device - Planning documented and maintained | - Design planning requirements established - Roles and responsibilities for design control activities defined - Planning scaled to device complexity/risk - Plan documentation requirements specified | - Design plan template with comprehensive content requirements - Named stages (Concept, Development, Verification, Validation, Transfer) with phase gates - Development matrix showing activities by stage - RACI matrix template with defined roles - Interface matrix including drug development for combination products - Traceability matrix requirements with tool specification - Competence matrix linkage - Risk management plan and usability engineering file integrated in planning | - Planning required but content/format not specified - Stages mentioned without clear definitions or criteria - Development planning vague or not mapped to stages - Responsibility assignment informal - Interfaces mentioned without management process - Traceability mentioned without methodology - Resources assumed without planning process - Plan updates ad hoc without triggers | - No design planning requirements - No defined design stages or phase gates - development not planned (treated as end-of-project only) - No responsibility assignment requirements - No interface management - No traceability requirements - No resource planning - Plans static with no update mechanism | - Agile/iterative planning options - Risk-based planning intensity by device class - Planning templates for common device types - Plan metrics and KPIs - Drug development interface for combination products - QMSR/820.30 alignment for US market
ISO7.3.3 | Design and Development Inputs | Determine whether design inputs are determined, documented, reviewed, and approved, including functional/performance/safety requirements, regulatory requirements, risk outputs, previous design information, and other essential requirements. | - Design input determination process established - Functional and performance requirements included - Safety requirements derived from risk management - Applicable regulatory requirements identified - Risk management outputs incorporated as inputs - Information from previous similar designs considered - Other essential requirements (manufacturing, packaging, service) addressed - Inputs reviewed for adequacy and approved - Incomplete/ambiguous/conflicting requirements resolved - [Combo] Drug product interface requirements as device inputs - [UE] User interface and use-related requirements included | - Input determination process defined with multiple sources - Functional, performance, and safety requirements explicitly required - Regulatory requirements and standards applicability assessed - Risk-derived requirements required as inputs - Previous design/predicate analysis required - Input review and approval process defined | - Input sources checklist (user needs, regulatory, standards, risk, predicate) - Requirement categorization (functional, performance, safety, regulatory) - Regulatory input checklist by market (US, EU, etc.) - Preliminary hazard analysis feeds design inputs - Predicate comparison template - Input review checklist with adequacy criteria - Drug-device interface specification for combination products - Task analysis and use scenario documentation for usability | - Inputs required but sources not comprehensive - Requirement categories not distinguished - Regulatory requirements informal or not specific to market - Risk linkage unclear or timing not specified - Previous design consideration ad hoc - Review process without adequacy criteria | - No input determination process - Functional/performance/safety requirements not distinguished - Regulatory requirements assumed, not documented - No risk-derived requirements - Previous designs not considered - No input review or approval | - Stakeholder needs analysis for input sources - Competitive device analysis - Medical/clinical advisory input process - Post-market feedback integration - Input freeze/baseline process - Drug CMC coordination for combination products - Use error analysis from predicates
ISO7.3.4 | Design and Development Outputs | Determine whether organization has established and maintained procedures which govern mandate the creation and control of design output records in verifiable form, approved before release, and include acceptance criteria, essential characteristics. | - Documented procedures for Design and Development Outputs exist - Procedure addresses the need for design outputs to meet the input requirements - Procedure establishes that Design Outputs shall be approved prior to release and that records shall be maintained - Procedure establishes that the outputs shall be in a form suitable for verification against inputs - Procedure establishes that Design Outputs shall contain appropriate information for purchasing, production and service provision - Procedure notes that Design Outputs shall contain or make reference to acceptance criteria - Procedure requires that Design Outputs specify the characteristics of the product that are essential for its safe and proper use. - Procedures maintained under document control - Scope and applicability clearly defined | - Design and development output procedure(s) documented and controlled - All 7.3.4 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled | - Comprehensive procedure with clear structure covering all 7.3.4 elements - Process flowchart providing visual overview - Defined terms and role definitions - Cross-references to related procedures (Design Inputs, Risk Management, Human Factors, Purchasing Controls) - Training requirements identified - Expectations established for when Design Outputs may be leveraged or included by reference (To vendor, partner) - Establishes what is in scope and out of scope between the combination product and drug product aspects (i.e. Drug Product quality specifications) - Discussion or inclusion by reference the need for a tracability matrix (Trace Matrix) to link User Needs, Design Inputs, Design Outputs and Verification/Validation activities. - Process for managing changes to Design Outputs is described or included by reference, establishing expectations for pre and post design freeze state. | - Procedures exist but incomplete coverage of 7.3.4 subclauses - Some elements in separate documents without clear linkage - Scope unclear or overly broad/narrow - Document control status uncertain - Over reliance on Drug Product or Drug Substance specification processes - No mention of how Design Outputs held by CMOs, Vendors or Partners will be managed or reviewed | - No documented design and development output procedure - Major gaps in 7.3.4 coverage (e.g., no expectation for product acceptance criteria or alignment with Design Inputs ) - Procedures not controlled documents - Procedures clearly inappropriate for device risk | - Clear identification of output types (e.g., drawings, specifications, software requirements, manufacturing instructions, labeling, packaging specs, risk control specifications) to prevent ambiguity. - Process and expectation that risk control measures (per ISO 14971) are translated into measurable, testable outputs and traceably documented. - Defined process for converting design outputs into supplier specifications and quality agreements to avoid misinterpretation. - Clear guidance on revision control, baseline creation, and change impact assessment once outputs are frozen. - Linkage to formal DFM/DFA review requirement to confirm outputs are feasible for production and inspection before release.
ISO7.3.5 | Design and Development Review | Determine whether organization has established procedures to ensure systematic design reviews are conducted at suitable stages with appropriate participants to evaluate design results, identify problems, and document actions. | - Documented procedures for Design and Development Reviews exist - Procedures maintained under document control - Scope and applicability clearly defined - Procedure establish suitable stages for review - Procedure establishes the need to both evaluate the ability of the results of the design and development effort to meet requirements as well as identify and propose necessary actions. - Procedure shall define minimum attendance/representation by function - Procedure shall establish that records of the review along with any actions shall be maintained along with details on the design under review, participants involved and the date of the review. | - Design and development review procedure(s) documented and controlled - All 7.3.5 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled | - Review types defined (technical, management, phase gate) - Stage-specific review criteria and checklists - Attendee matrix by review type with quorum requirements - Review minutes template with action tracking - Training requirements identified | - Reviews required but stages unclear - Problems addressed informally - Participation required but functions not specified - Independent reviewer expectations unclear - Procedures exist but incomplete coverage of 7.3.5 subclauses - Some elements in separate documents without clear linkage - Scope unclear or overly broad/narrow - Document control status uncertain | - No documented design and development review procedure - Major gaps in 7.3.5 coverage (e.g., no expectation for action tracking or ensuring representative functions in attendance) - Procedures not controlled documents - Procedures clearly inappropriate for device risk | - Clear articulation of what each phase review must confirm (e.g., feasibility, risk control adequacy, verification readiness, transfer readiness). - Predefined requirements for what must be completed before the review can occur and what constitutes successful closure. - Explicit requirement to review the current risk file status, design development plan and Design History File Index - Defined requirement for review materials to be circulated in advance to improve review quality. - Action item closure tracking system or process - CAPA linkage process or guidance for systemic issues - Expectation for key design decisions, failures or technical issues to be discussed and documentation - Expectations around phase appropriate attendance from cross functions (Drug development, Clinical, Commercial, Etc) - Differentiation between Phase reviews and technical reviews - Process for determining when an independent reviewer is needed and selection process - Phase appropriate review questions and key topics
ISO7.3.6 | Design and Development Verification | Determine whether design verification is performed per planned arrangements to ensure outputs meet inputs, with appropriate records, risk control verification, and formative usability evaluation. | - Verification per planned arrangements - Output conformance to inputs demonstrated - Verification methods appropriate - Records of results and actions maintained - Risk control implementation and effectiveness verified - [Combo] Verification with drug product or surrogate | - Verification per planning requirements - Output-input conformance required with traceability - Method selection and validation (test, analysis, inspection) defined - Protocol and report requirements - Risk control verification required (implementation and effectiveness) - Drug product or surrogate use in verification for combination products | - Verification matrix showing methods per output - Protocol template with statistical rationale requirements - 100% input coverage requirement with gap analysis - Verification report template with deviation handling - Risk control verification matrix with effectiveness criteria - Surrogate/bridging rationale for drug product testing | - Verification required but planning link unclear - Conformance mentioned without method - Protocol format undefined - Records incomplete - Risk control verification vague - Drug use mentioned without criteria | - No verification per plan - Inputs-outputs disconnected - Methods ad hoc or not validated - No verification records - Risk controls not verified - Device verified without drug product or appropriate surrogate | - Verification vs validation distinction guidance - Sample size determination methodology - Worst-case condition selection - Accelerated testing guidance - Negative/boundary testing
ISO7.3.7 | Design and Development Validation | Determine whether design validation is performed per planned arrangements to demonstrate conformance to user needs and intended uses, completed prior to release, with summative usability validation and overall residual risk evaluation. | - Validation per planned arrangements - Conformance to user needs and intended uses demonstrated - Actual or simulated use conditions - Production units or equivalents used - Completed prior to product release - Records of results and actions maintained - Summative usability validation performed - Overall residual risk evaluation completed - [Combo] Validation with commercial-equivalent drug product or surrogate | - Validation per planning requirements - User needs traceability demonstrated - Use conditions defined (actual or simulated with justification) - Production equivalence criteria specified - Timing before release enforced - Validation records requirements - Summative HF study requirements - Overall residual risk evaluation before release - Drug product or surrogate requirements for combination products | - Validation matrix showing user needs coverage - Intended use acceptance criteria - Use environment specification - Production equivalence checklist - Validation as release gate criterion - Validation report template with statistical analysis - Summative protocol with 15+ users per group, critical task criteria - Risk management report with risk-benefit methodology - Clinical protocol integration for combination products | - Validation required but planning link vague - User needs link informal - Use conditions unclear - Production equivalence not defined - Timing not enforced - Records incomplete - Summative criteria undefined - Residual risk evaluation informal - Drug or surrogate use without requirements | - No validation per plan - User needs not validated - Inappropriate use conditions - Prototype acceptable for validation - Validation after release acceptable - No validation records - No summative study - No overall residual risk evaluation - Device-only validation for combination product | - Clinical study integration - Real-world evidence planning - Patient vs HCP validation distinction - Home vs clinical use validation - Training effectiveness validation - Post-market validation updates - Long-term use validation
ISO7.3.8 | Design and Development Transfer | Determine whether procedures ensure design outputs are correctly translated to production specifications with verified manufacturing suitability and capability. | - Transfer procedures documented - Design outputs translated to production specifications - Manufacturing suitability verified - Production capability meets product requirements - Process validation linkage - Labeling and training materials finalized and released | - Transfer procedure requirements defined - Translation process to manufacturing specs - Manufacturing suitability required - Capability demonstration required - Process validation linkage - Labeling/training release at transfer | - Transfer checklist - DMR completeness verification - Capability study requirements (Cpk/Ppk criteria) - Equipment qualification and operator training - Process validation protocol linkage - First article inspection requirements - Labeling approval process with multi-language requirements | - Transfer mentioned without procedure details - Translation process vague - Manufacturing suitability verification informal or not documented - Capability criteria undefined - Process validation disconnected - Labeling timing unclear | - No transfer procedures - Design-manufacturing gap - No suitability verification - Capability not demonstrated - No process validation link - Labeling/training uncontrolled | - Supplier transfer requirements - Training documentation transfer - Scale-up considerations - Multi-site transfer - Continuous verification requirements - Label printing qualification
ISO7.3.9 | Control of Design and Development Changes | Determine whether design changes are identified, reviewed, verified/validated as appropriate, approved before implementation, and assessed for impact on constituent parts, delivered product, risk, and human factors. | - Changes identified and records maintained - Changes reviewed, verified, validated as appropriate - Changes approved before implementation - Impact on constituent parts evaluated - Impact on delivered product assessed - Risk file update assessed - HF impact assessed - [Combo] Drug-device interface change impact assessed | - Change identification and documentation process - V&V determination process for changes - Approval workflow before implementation - Constituent part impact assessment - Delivered product/field action assessment - Risk file review for changes - HF impact assessment - Drug product impact for combination products | - Change request form/template - Change classification criteria (minor, major, critical) - V&V matrix for change types - Approval authority matrix - Impact assessment template including regulatory filing impact - Field action/recall determination criteria - Risk change impact matrix with new hazard identification - HF revalidation decision tree - Drug-device change matrix with supplement determination | - Changes addressed but process incomplete - V&V determination unclear - Approval informal - Impact assessment vague - Delivered product not addressed - Risk assessment mentioned without process - HF impact informal - Drug impact unclear | - No change control requirements - No V&V for changes - Changes implemented without approval - No impact assessment - Delivered product ignored - Risk file not updated - HF changes uncontrolled - Drug-device impact not assessed | - Expedited change process for safety - Risk-based change classification - Change metrics and trending - Batch/lot impact assessment - Supplier change notification process - Post-market surveillance linkage - FDA/EMA coordination for combination products
ISO7.3.10 | Design and Development Files | Determine whether the procedure defines the requirement for maintenance of a design and development file for each device type/family containing conformity records and change records for the product under development. | - Documented procedures for the creation and maintenance of the Design and Development File (aka Design History File) exist - Procedures establish the need to both create and maintain for each medical device type, combination product type or product family - Procedure establishes that the file shall include or make reference to applicable records generated during development - Procedure estalishes the need for the records associated with design and development changes | - Design and development file requirement documented and controlled (in standlone procedure or by reference in another procedure) - All 7.3.10 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled | - Comprehensive procedure with clear structure covering all 7.3.10 elements - Template which establishes expected format or process of documenting content - Defined terms and role definitions - Cross-references to related procedures (Risk Management, Human Factors) - Training requirements identified - Procedure contains instruction on how to make reference to records held outside of company | - Procedure or section of procedure exists but incomplete coverage of 7.3.10 subclauses - Some elements in separate documents without clear linkage - Scope unclear or overly broad/narrow - Document control status uncertain - File required but structure or templates unclear - Type/family criteria vague - Conformity records incomplete - No guidance on how to reference records external to company - No expectations established for when file is updated and reviewed (i.e. at Design Reviews, etc) | - No documented requirement for the creation/maintenance of a development file - Major gaps in 7.3.10 coverage - Procedures clearly inappropriate for device risk - Procedure confuses document control system with Development File | - Templates provided (if not electronically established) - Guidance provided on making reference to external files and documents - Expectations established for review and update cadence (ever document revision, at phase reviews, etc) - Process established for post approval/launch maintenance of file - Linkages to Change Management procedure and corporate level policies established - Expectations and guidance established for what doesnt belong in the Development File (Drug development documents, etc)

## Manual Verification
ID/Clause Number | ID/Clause Title | Minimum Subset to Pass Bullets Add as many as appropriate | Client 1 | Client 2 | Client 3 | Client 4
ISO7.3.1 | General Requirements | - Design and development procedure(s) documented and controlled - All 7.3 subclauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled |  |  |  |  | Meets requirement
ISO7.3.2 | Design and Development Planning | - Design planning requirements established - Roles and responsibilities for design control activities defined - Planning scaled to device complexity/risk - Plan documentation requirements specified |  |  |  |  | OFI
ISO7.3.3 | Design and Development Inputs | - Input determination process defined with multiple sources - Functional, performance, and safety requirements explicitly required - Regulatory requirements and standards applicability assessed - Risk-derived requirements required as inputs - Previous design/predicate analysis required - Input review and approval process defined |  |  |  |  | Deficiency
ISO7.3.4 | Design and Development Outputs | - Design and development output procedure(s) documented and controlled - All 7.3.4 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled
ISO7.3.5 | Design and Development Review | - Design and development review procedure(s) documented and controlled - All 7.3.5 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled
ISO7.3.6 | Design and Development Verification | - Verification per planning requirements - Output-input conformance required with traceability - Method selection and validation (test, analysis, inspection) defined - Protocol and report requirements - Risk control verification required (implementation and effectiveness) - Drug product or surrogate use in verification for combination products
ISO7.3.7 | Design and Development Validation | - Validation per planning requirements - User needs traceability demonstrated - Use conditions defined (actual or simulated with justification) - Production equivalence criteria specified - Timing before release enforced - Validation records requirements - Summative HF study requirements - Overall residual risk evaluation before release - Drug product or surrogate requirements for combination products
ISO7.3.8 | Design and Development Transfer | - Transfer procedure requirements defined - Translation process to manufacturing specs - Manufacturing suitability required - Capability demonstration required - Process validation linkage - Labeling/training release at transfer
ISO7.3.9 | Control of Design and Development Changes | - Change identification and documentation process - V&V determination process for changes - Approval workflow before implementation - Constituent part impact assessment - Delivered product/field action assessment - Risk file review for changes - HF impact assessment - Drug product impact for combination products
ISO7.3.10 | Design and Development Files | - Design and development file requirement documented and controlled (in standlone procedure or by reference in another procedure) - All 7.3.10 clauses addressed (directly or by reference to other procedures) - Scope defines applicable products/projects - Procedure approved and revision controlled